## Slide 1: ILS Characterization ExoMars TGO NOMAD Solar Occultation (SO) channel Giuliano Liuzzi Geronimo Villanueva NASA/GSFC – March-May/2021

## Slide 2: Method to derive the ILS: self-correlation

## Slide 3: Method to derive the ILS: self-correlation
This technique uses the intrinsic information content of the spectrum y(x) in a completely agnostic way, i.e. not assuming any a-priori lineshape functional form, nor any hypothesis about how well isolated lines are.
The method relies only on a good high-resolution model R(x) for the signal, which must include a reliable model for the AOTF in order to include all visible lines.
If the ILS varies locally within the spectrum, the ILS can be calculated on narrow spectral interval. However, this will introduce more noise, so in deriving the ILS a trade-off needs to be established.

## Slide 5: Order 165: observation (10x) vs. model (RP 200,000)
H = 130 km, bin 3
30% of the spectrum
Shift = +/- 1.2 cm-1
j changes

## Slide 6: Order 165: functional form for the ILS
H = 130 km, bin 3

## Slide 7: Order 165: functional form for the ILS
ILS computed from the spectrum
Modeled ILS

## Slide 8: Order 165: are we immune from altitude (i.e. line saturation)?
b0…3 = 1.17410e-09, -2.81659e-06, 1.68969e-03, -2.27274e-03
(shift -0.002 to 0.288 cm-1)
= 9.30255e-02 cm-1
A2 = 2.36054e-01
b0…3 = -1.35561e-10, -2.47316e-06, 1.50305e-03, 7.37273e-03
(shift 0.007 to 0.23 cm-1)
s = 1.05804e-01 cm-1
A2 = 2.85389e-01
H = 130 km, bin 3 (not saturated)
H = 70 km, bin 3 (VERY saturated)
b0…3 = -2.15478e-09, -3.44165e-07, 1.20833e-03, 9.42663e-03
(shift 0.009 to 0.289 cm-1)
= 9.78230e-02 cm-1
A2 = 2.36054e-01
H = 100 km, bin 3 (partially saturated)

## Slide 9: Current recipes: a comparison

## Slide 10: * Applied to both gaussians
 | A1 | A2 | s1 | s2 | b1 | b2 | Tot. pars
BIRA | 2 | 2 | 2 | 2 | 2 | 2 | 12
IAA |  | 2 | 2 |  | 2* | 2 | 8
GSFC |  | 1 | 1 |  |  | 3 | 5

## Slide 11: * Applied to both gaussians
 | A1 | A2 | s1 | s2 | b1 | b2 | Tot. pars
BIRA | 2 | 2 | 2 | 2 | 2 | 2 | 12
IAA |  | 2 | 2 |  | 2* | 2 | 8
GSFC |  | 1 | 1 |  |  | 3 | 5
The width is defined by the grating grooves and its quality. There is only one grating, s1≡s2 ≈ s
1
1

## Slide 12: * Applied to both gaussians
 | A1 | A2 | s1 | s2 | b1 | b2 | Tot. pars
BIRA | 2 | 2 | 2 | 2 | 2 | 2 | 12
IAA |  | 2 | 2 |  | 2* | 2 | 8
GSFC |  | 1 | 1 |  |  | 3 | 5
Shifts to the main kernel imply issues with the frequency calibration, not the ILS
1
1
The width is defined by the grating grooves and its quality. There is only one grating, s1≡s2 ≈ s

## Slide 13: * Applied to both gaussians
 | A1 | A2 | s1 | s2 | b1 | b2 | Tot. pars
BIRA | 2 | 2 | 2 | 2 | 2 | 2 | 12
IAA |  | 2 | 2 |  | 2* | 2 | 8
GSFC |  | 1 | 1 |  |  | 3 | 5
In half of the detector, the two gaussians overlap, therefore it is hard to accurately derive A1/A2
1
1
1
1
Shifts to the main kernel imply issues with the frequency calibration, not the ILS
The width is defined by the grating grooves and its quality. There is only one grating, s1≡s2 ≈ s

## Slide 14: Aperture
AOTF
Slit
Parabolic mirror
GRATING
Detector
1
2

## Slide 15: Gaussian amplitudes: and

## Slide 16: Gaussian widths: and

## Slide 17: Gaussians’ shift:-

## Slide 18: 2-D images of the solutions

## Slide 19
GSFC old
GSFC new
BIRA
IAA

## Slide 20: Comparison between fits using different solutions

## Slide 21: Order 165: spectra between 65 and 160 km fit
To test the different solutions, we have created a Python code that collects them into a module, which is then used to convolve high-res radiances generated by PSG. The Python code also uses the AOTF solution as provided by Geronimo in the new update.
The code fits spectra from order 165, bin 3 from a full scan with all the ILS solutions and compares both retrieved CO2 scaler (w.r.t. GEM) and the chi-square.
This is a way to directly compare the quality of the solutions with the data, and to iterate to further refine the ILS parameterization. It is also a way to understand how many parameters we effectively need to produce a meaningful ILS solution.

## Slide 22: H = 67 km

## Slide 23: H = 85 km

## Slide 24: H = 103 km

## Slide 25: H = 113 km

## Slide 26: H = 131 km

## Slide 27: H = 141 km

## Slide 28: H = 160 km

## Slide 29
All the solutions look consistent, with differences only in the CO2 scaler around 100 km (lines start to
“de-saturate” and ILS is critical)
BIRA
IAA
GSFC-old GSFC-new
BIRA
IAA
GSFC-old GSFC-new

## Slide 30: Overall recipe

## Slide 31
Using the same methodology for all the orders, we have first derived for each of them the s, A2 and the polynomial shift of the two gaussians.
Then we fit the average polynomial shift and derive its final expression. The s, A2 are constant across the orders and calculated as the average of the values obtained by each order. We could not identify any specific trend of these parameters with diffraction order (more on that in a few slides).

## Slide 32: Gaussians’ shift:- for several orders
Pixel number
Wavenumber x 3700/vcent

## Slide 33: Is there s and/or variability across orders and pixels?
Cross-correlation is not able to provide solid constraints on these two parameters, which are better constrained by searching for the best fit quality, varying them.
For each order, we do the fitting of data at non-saturated altitudes for several values of s and A2, which allows to see if there is any variability with order and pixel.
Resolving power (~ s)
Amplitude A2
Example: order 165, H=131 km

## Slide 34
Example: order 165, H=131 km
Pixels 0 to 100
Pixels 240 to 320

## Slide 35
Example: order 141, H=50 km
Pixels 0 to 100
Pixels 240 to 320

## Slide 36
Example: order 122, H=35 km
Pixels 0 to 100
Pixels 240 to 320

## Slide 38: CO2 profiles using different orders

## Slide 39: ILS GSFC new solution

## Slide 40
Every profile is characterized by a different level of opacity. We thought about possible hypotheses to explain those and the sharp minima in the profiles at the altitude where lines start to become less saturated:
Bias in the dark correction
ILS issues
AOTF offset/gaussian continuum

## Slide 42: ILS GSFC new solution + 25% offset